## Slide 1: Transradial Prosthetic Arm
Kranti Peddada
Team Members: Kendall Gretsch and Henry Lather
Mentors: Dr. Charles Goldfarb and Dr. Lindley Wall
Website: www.transradialprostheticarm.weebly.com

## Slide 2: Abbreviated Design Specifications
Patient Population
  Unilateral, transradial limb difference
  Ages 5+
Total Parts Cost
  Maximum $150
Functionality
  Hand and thumb open and close at mouth, belt, and in front
  Independent thumb movement with key grip
  Ability to lift and hold at least 500 g
Weight
  Does not exceed weight of missing limb

## Slide 3: New Design Specifications
Do-it-yourself, open-source manufacturing and assembly
Robohand Project distribution model
  Primarily through online support groups
Advantages
  Rapid distribution
  Wide audience
  Low cost
  Highly customizable
  Avoids hassles of regulations
Kept in mind when choosing component of final design

## Slide 4: Overview of Design
Inertial Measurement Unit (IMU)
Velcro strap
Wire
Battery Pack

## Slide 5: Overview of Design

## Slide 6: Analytic Efforts

## Slide 7: IMU and Sampling Rate
IMU from Prof. James Bobrow of UC Irvine
Internal update rate of 250 Hz using SPI transmission
No disadvantages in using maximum sampling rate
250 Hz chosen
2.54 cm

## Slide 8: Power Requirements
IMU = 72 mA
Arduino Micro = 28 mA
Motors at rest = 6.4 mA (x5)
Motors at peak torque and angular velocity = 428 mA (x5)
At rest ≈ 132 mA
At peak ≈ 2.2 A
Chosen voltage regulator can handle constant current of 3 A (maximum of 5 A)

## Slide 9: Torque Requirements
Our motors can supply up to 2.2 kg-cm

## Slide 10: Specific Details

## Slide 11: IMU Operation
Outputs yaw, pitch, roll, their time derivatives, and linear acceleration in x, y, and z axes

## Slide 13: Up/Down Shoulder Movement

## Slide 14: Fwd./Rev. Shoulder Movement

## Slide 15: Software Flow Chart
Initialization:
Roll angle read
Device ON
Read IMU Data
IF:
Hand closed
Thumb closed
Shoulder up
IF:
Hand open
Thumb closed
Shoulder down
IF:
Hand closed
Thumb closed
Shoulder forward
IF:
Hand closed
Thumb open
Shoulder backward
Hand OPEN
Hand CLOSED
Thumb OPEN
Thumb CLOSED

## Slide 16: Hand

## Slide 17: Socket

## Slide 18: Total Assembly

## Slide 19: Safety
No major safety concerns anticipated
  Only DC current used
  All materials are non-toxic
  NiMH batteries are safe
  Motors kept distal from residual limb

## Slide 20: Costs
Item | Quantity (# or # pkgs.) | Total Cost
Micro servo motor | 5 | $ 46.49
Electrical tape | 1 | $ 0.76
Fishing line | 1 | $ 2.51
Elastic cord (stringing) | 1 | $ 1.63
ABS filament | n/a | $ 6.20
Superglue | 1 | $ 2.49
#6 lock washers | 12 | $ 1.29
#6-32 cap nuts | 2 | $ 2.59
#8 3/8" set screws | 6 | $ 2.33
#8-32 1/2" machine screws | 12 | $ 2.59
#8-32 screw nuts | 8 | $ 2.59
#6-32 6" threaded rod | 1 | $ 3.06
#216-1/2" screw eye | 4 | $ 1.29
Adhesive wrap | 1 | $ 3.75
Total |  | $ 226.12
Item | Quantity (# or # pkgs.) | Total Cost
Inertial measurement unit | 1 | $ 24.99
Velcro straps | 1 | $ 4.15
Stranded 18 AWG wire | 1 | $ 9.24
Printed circuit board | 1 | $ 2.43
Heat shrink tubing | 1 | $ 2.28
AA battery holder | 1 | $ 7.51
AA NiMH rechargeable batteries | 8 | $ 30.88
AA battery charger | 2 | $ 26.77
5V switching voltage regulator | 1 | $ 12.55

## Slide 21: 3D Printing and Assembly
All parts are off-the-shelf or can be shipped in 2-5 business days
Socket and hand are 3D printed
  Less than 15 hours of printing time
Basic soldering experience required
Basic proficiency with hand-held electric tools required
Fitting should be overseen by a Certified Prosthetist

## Slide 22: The Prototype

## Slide 23: Evaluation of Prototype
Patient Population
  Unilateral, transradial limb difference
  Ages 5+
Total Parts Cost
  $226.12
Functionality
  Hand and thumb open and close at mouth, belt, and in front
  Independent thumb movement with key grip
  Ability to lift and hold 50 g (still testing)
Weight
  Weighs 240 g

## Slide 24: Intellectual Property
Yes!
  New, useful, and non-obvious
However, we want our design to be open-source
  Creative Commons License
  GNU General Public License

## Slide 25: What We Learned
Specific Skills
  Henry: Arduino boards, Arduino IDE, SPI transmission, PWM motor control, power supplies, general embedded systems
  Kendall: prototyping mechanical systems, Autodesk Inventor, using electric power tools
  Kranti: theory of servo and stepper motor control, micro-controller design and theory
Keep design notebooks up-to-date
Use Pugh charts to choose design
1 electrical engineer + 1 mechanical engineer + 1 computer engineer > 3 biomedical engineers

## Slide 26: Future Directions
Continue prototyping
Deliver finished prototype to Dr. Goldfarb and Dr. Wall
Find suitable candidate to further test device
Make adjustments as needed
Give candidate the device
Publish all designs and instructions to the internet

## Slide 27: Acknowledgements
Prof. Humberto Gonzalez
Patricia Widder
Our mentors:
  Dr. Charles Goldfarb
  Dr. Lindley Wall

## Slide 28: Thank You
Questions?